## Slide 1: Exercises
Dr. Ronald Hesse

## Slide 2
General structure of the following exercises
  Description and aim of the exercise,
    e.g. Exercise 1: Programme handling.
  Demonstration of the task.
  Step-by-step guide to perform the given task,
  e.g. Task 1: Programme handling.
  You are kindly asked to attentively follow the demonstrations first.
  Afterwards there will be enough time to process the tasks yourself.
  In case of questions, please ask the assistants.
  With progress of the course less instructions are given
  and more self-responsible work is necessary.
  The results of the tasks are saved in ‘Exercise-xx-Result.ufp’.

## Slide 3
1. Structure and concept of the programme
Exercise 1
First steps – getting to know the programme.
Setting of programme parameters.
Show and hide standard windows.
Show all spectrum-window types.
Handling ‘Preferences’ and pull-down and pop-up commands:
  setting ‘Preferences’ without an opened spectrum window (e.g. language, toolbar, modify toolbar)
  setting ‘Preferences’ with an opened project (e.g. dialog font/size, axis colour, axis labelling, loading and saving of presettings…).
Information in headline of the windows.
Overview over different programme parts (e.g. modify, batch processing, concentration, information).
Exercise-01-Start Presentation of all 14 spectrum-window types generated by UNIFIT

## Slide 4
1. Structure and concept of the programme
Task 1
Start the programme UNIFIT 2024 and open the dialogue ‘Preferences’ with the pull-down command [Preferences → Preferences]. Nine groups of preferences are defined. Check the settings of every group.
General: a) Modify the general programme parameters with ‘Modify’. Define ‘Number of the First Standard Window’ to: 41. b) Change the language from English to German and back. c) Deactivate and activate the Toolbar. Customize the toolbar with ’Modify’. d) Change the text of the dialogue. Increase and decrease the text size with ‘Alt +’ and ‘Alt -’
Display: Display the design preferences by call the design dialogues: a) ‘Graphs Standard Window/Wagner Plot’. b) ‘Axes/Lines/Text’ c) ‘3D Plot Waterfall/XY 3D Plot Colour Profile’ d) ‘3D Plot Waterfall 0°Plus’ e) ‘Parameter Plot’

## Slide 5
1. Structure and concept of the programme
Task 1
Sources Check the definition of the excitation satellites with ‘Define Sources’.
Close the dialogue and open the test project ‘Exercise-01.ufp’ with [File → Open Project] in the folder ‘Documents\Unifit_2024_User_Files\Exercises’.
Transfer the settings of the spectrum windows of the loaded project to the preferences with [Preferences – Transfer Project Settings to Preferences].
Activate step by step every spectrum window, recognize the differences of the window types and use the different pop-up commands for the modification of the spectrum design: Window 1: Wagner Plot of Ag compounds Window 2: Plot 3D Colour Profile Window 3: Plot 3D Waterfall 0° Window 4: Plot 3D Waterfall 45° Window 5: Plot 3D Waterfall -45° Window 6: Parameter Plot of an angle resolved measurement

## Slide 6
1. Structure and concept of the programme
Task 1
Window 7: Plot 3D Waterfall 0°Plus Window 8: XY 3D Colour Profile Window 9: XY 3D 45° Colour Profile Window 10: XY 3D Plot 45° Window 11: XY 3D -45°Colour Profile Window 12: XY 3D Plot -45° Window 13: Image Window 42: Standard Window of a fitted As2p3 spectrum of GaAs Window 46: Standard Window of a survey of GaAs
Transfer separate window settings to preferences.
Open the dialogue ‘Preferences’. Load the saved presetting file ‘Presetting_Image_Export.SET’.
Transfer the new settings to the graphs and coordinate-system design of the different windows.

## Slide 7
1. Structure and concept of the programme
Task 1
Activate window 46, open the dialogue ‘Preferences’ with [Preferences → Preferences] and change the setting of the X-Axis and Y-Axis’.
Call [File → Original/Accept Preferences Active Window] and practice [Modify
  → Charge correction],
  → Differentiation],
  → Reduction], (after setting boundaries) and expansion of the spectrum,
  → Spike Correction], set up new data points by pressing the left mouse key, deactivate spike correction,
  → Spectrum Manipulation → Energy Shift],
  → Spectrum Manipulation → Correct Intensity],
  → Normalization].
Recreate the original spectrum by [File → Original...] and show spectra information via [Information/Editing → Acquisition Parameters] or by pressing the middle mouse key.

## Slide 8
2. Background subtraction and peak fit
Exercise 2
Fit of the Shirley background function before or parallel to the peak fit.
Peak fit with different model functions (sum, product or convolution of a Lorentzian and a Gaussian function).
Input of start parameters manually or by loading saved parameters.
Using relative fit parameters with more than one master peak.
Exercise-02A-Start one peak, Voigt profile, peak intensity: 1000 Counts, position: 190 eV, Lorentzian width: 1.7 eV, Gaussian width: 0.3 eV
Exercise-02B-Start one peak, Voigt profile, peak intensity: 1000 Counts, position: 190 eV, Lorentzian width: 1.0 eV, Gaussian width: 1.0 eV
Exercise-02C-Start five peaks, sum, two master peaks (peak 1 and peak 3), L/G mixing: 0.5, FWHM: 2 eV, asymmetry: 0, two peak groups: 1. group: (peak 1 and peak 2): intensity ratio: 0.5, peak separation: 2 eV, intensity peak 1: 800 counts, peak position peak 1: 188 eV 2. group: (peak 3, peak 4 and peak 5): intensity ratio peak 4/peak 3: 0.666, intensity ratio peak 5/peak 3: 0.333, peak separation: peak 4 – peak 3: 3 eV, peak 5- peak 3: 6 eV, intensity peak 3: 600 counts, peak position peak 3: 189 eV

## Slide 9
2. Background subtraction and peak fit
Task 2A Comparison of background subtraction before and parallel to the fit
Change in the dialogue ‘Preferences’ the fit procedure to ‘Convolution‘, ‘Absolute’ and ‘Homogeneous Sample’.
Open test spectrum ‘Exercise-02A-Start.ufp’ and duplicate the spectrum via the pop-up command ‘Programme Internal Coping - Programme Internal Insertion’.
First Window
Calculate an appropriate background with [Modify → Calculate Background → Constant] and subtract this background by pressing ‘Subtract’.
Manually set up the peak fit with [Peak Fit → Manual Input of Start Parameters → Singlet Peak/Convolution/Absolute] for one component or with the pop-up command ‘Peak Fit Input Parameters Singlet Peaks’ and run the peak fit.
Second Window
Run the peak fit with a fittable background via [Peak Fit → Fit Background XPS] or with the pop-up command ‘Fit Background XPS (HOM)’ and one peak.
Compare residual functions of both spectra.

## Slide 10
2. Background subtraction and peak fit
Task 2B Fit of a Voigt test spectrum with different model functions
Open test spectrum ‘Exercise-02B-Start.ufp’.
Perform the peak fit in three windows with three different model functions [Preferences – Preferences] Fit Procedures → Sum/Product/Convolution and fittable background (Shirley + Polynomial).
Compare residual the functions, χ2* and Abbe criteria.

## Slide 11
2. Background subtraction and peak fit
Task 2C Comparison of different numbers of master peaks
Change the fit procedure to ‘Sum‘ and ‘Relative’.
Open test spectrum ‘Exercise-02C-Start.ufp’ and duplicate the spectrum.
First Window
Load the fit parameters ‘Exercises\Exercise-02C-OneMasterPeak.par’ with [Peak Fit → Load Start Parameters...[.
Confirm the fittable background with ‘OK’, confirm the Min/Max-values with ‘OK’, confirm the fit parameters with ‘OK’ and start the iteration (number of iterations: 22, number of cycles: 2).
Second Window
Load the fit parameters ‘Exercises\Exercise-02C-TwoMasterPeaks.par’.
Confirm the loaded fit parameters and start the iteration.
Compare the results of the peak fit and the residual functions of both spectra.

## Slide 12
3. Peak fit and calculation of the fit result errors
Exercise 3
Influence of fit parameter on the peak fit.
Calculation of the fit parameter errors.
Minimization of the fit parameter errors by elimination of redundant fit parameters (fit parameters with an error >100% need to be set to zero and to be fixed).
Exercise-03 -Start two peaks, Voigt profile, background polynomial, statistic noise, 1. 10,000 counts, Lorentzian width 1.7 eV, Gaussian width: 0.3 eV, position: 192 eV, 2. 50,000 counts, Lorentzian width: 0.3 eV, Gaussian width: 1.7 eV, position: 190 eV

## Slide 13
3. Peak fit and calculation of the fit result errors
Task 3
Change setting of the peak fit to ‘Convolution’ and ‘Absolute’ and error calculation to ‘Matrix Inversion’.
Load project ‘Exercise-03-Start.ufp’.
Choose a fittable background with Shirley and constant share.
(a, e free, B = 0 fix, Tougaard: C = 1643 fix, C’ = 1 fix, D = 0 fix, T0 = 0, fix).
Manual input of two components, perform the peak fit, start with 44 iterations.
Correct the fittable background appropriately
(additional polynomial, e.g. b = 1 variable, c = 0,1 variable).
If the fit was successful, calculate the errors of the fit parameters via [Peak Fit - Show Fit-Parameter Errors]. Display the relative uncertainties.
Choose parameters having an error >100 %, set them to zero and fix them.
Repeat the peak fit.
Repeat the error calculation.
Compare the fit-parameter errors of both results.

## Slide 14
4. Coupling of fit parameters for the peak fit
Exercise 4A Fitting with coupled doublet parameters
Performance of a satellite subtraction.
Usage of doublet lines and coupled fit parameters.
Peak fit with and without relative fixing of fit parameters.
Appropriate setting of fixed parameters.
Note: Even though the general fit procedure is ‘Absolute’, the second peak of a doublet is always treated relative to the first one.
Exercise-04A-Start real spectra for the analysis of the chemical states of sulfur, oxygen, nitrogen and carbon

## Slide 15
4. Coupling of fit parameters for the peak fit
Task 4A Fitting with coupled doublet parameters
Change preferences to ‘Sum’, ‘Absolute’ and ‘Matrix inversion’.
Load the project ‘Exercise-04A-Start.ufp’ and activate the sulfur peak and duplicate the spectrum via the pop-up command ‘Programme Internal Coping - Programme Internal Insertion’.
Operate in both windows: a) Perform the satellite subtraction with [Modify - Subtract Satellite]. b) Generate a fittable background (a, b, c, e free, d, Tougaard: B = 0, fix).
First Window
Perform the peak fit of the Sulphur spectrum with: all fit parameters free (asymmetry parameters 0, fixed), 44 iterations.
Second Window
Perform the peak fit of the Sulphur spectrum with: couple all doublet values appropriately (asymmetry parameters 0, fixed), 44 iterations.
Try the calculation of the uncertainties of the peak fit of both windows.
Compare the fit results and the fit-parameter uncertainties.

## Slide 16
4. Coupling of fit parameters for the peak fit
Exercise 4B Fitting with relative parameters
Usage of singlet peaks and relative fit parameters.
Appropriate fixing of the parameters.
Fitting with relative parameters.
Estimation of an energy loss function.
Exercise-04B-Start real data for the analysis of the oxidation of GaAs by ozone

## Slide 17
4. Coupling of fit parameters for the peak fit
Task 4B Fitting with relative parameters
Change peak-fit setting to ‘Sum’ and ‘Relative’.
Open project ‘Exercise-04B-Start.ufp’ and activate the As 2p3/2 spectrum.
Fit the As 2p3/2 peaks with a fittable background (a, b, c, e free, d fix, Tougaard: B, C, C’, D, T0 fix) and five singlets: GaAs 1323 eV As + 1.2 eV
AsO + 2.6 eV
As2O3 + 3.3 eV
As2O5 + 4.3 eV
Fix the relative fit parameters of the peaks 2 to 5 appropriately e.g. L-G Mixing = 1, fix
FWHM = 1, fix.
Run the peak fit with 22 iterations.
Label the lines of the five chemical species with [Information/Editing - Identify Lines], the dialogue ‘XPS/AES Lines’ appears, press the left mouse button and move the mouse over the spectrum, select As 2p3. ‘Select Line Type’: Chemical Shift, annotate the chemical components with ‘Annotation’.
Display the fit parameter table via ‘Peak Fit → Show/Correct Fit Parameters…’.

## Slide 18
4. Coupling of fit parameters for the peak fit
Exercise 4C Relative fit parameters and improved Tougaard background
Peak fit of a wide range spectrum.
Usage of singlet peaks, relative fit parameters and Tougaard background.
Manual definition of Tougaard background parameters.
Influence of the Tougaard background parameter T0 on the fit.
Parallel calculation of the loss function integrated into the Tougaard background.
Plot of loss functions in separate windows and one window.
Estimation of the energy gaps.
Plot of the inelastic electron scattering cross section functions.
Exercise-04C-Start wide range spectrum of SiO2

## Slide 19
4. Coupling of fit parameters for the peak fit
Task 4C Relative fit parameters and improved Tougaard background
Set the peak-fit settings to ‘Convolution’ and ‘Relative’.
Load the project ‘Exercise-04C-Start.ufp’.
Load the transmission function ‘ESCALAB220_MONO_LAXL_50EP.trm’ via [Preferences → Preferences], ‘Load T(E)’, The loaded transmission function is displayed..
Correct the survey spectrum with the transmission function via [Modify → Correction with IERF: ESCALAB220_MONO_LAXL_50EP.trm].
Reduce the spectral range appropriately (e.g. 235 eV – 84 eV).
Duplicate the spectrum.
First Window
Manual alignment of B, C, C‘, D and T0 with [Modify → Calculate Background →Tougaard], recommended values e.g. B = 450
C = 370 D = 680
C‘ = -1.3 T0 = 0 – 8 eV Adjust B-Parameter: Unchecked
Modify T0 and press ‘Preview’ to see the changes and save the manually defined loss function in ‘loss.cro’.

## Slide 20
4. Coupling of fit parameters for the peak fit
Task 4C Relative fit parameters and improved Tougaard background
Activate the fittable background calculation and load the saved loss function ‘loss.cro’ (settings: a, b, c free, d = 0, e = 0, fix (Shirley), Tougaard: B, C, C’ D free, T0 = 0 fix)
Set up two lines with [Peak Fit → Manual Input of Start Parameters → Singlet Peaks] or use the corresponding pop-up command.
Set and fix GP-FWHM of the second peak to 1.
Run the peak fit, use 22 iterations. Save the loss function with a suitable name.
Second Window
Repeat the peak fit with T0 = 4, free.
Save the loss function with a suitable name.
Compare the peak fit quantities.
Load and compare both loss functions with [Information/Editing → Plot Cross Section…].
Display both loss functions in one window by ‘Batch Processing → Plot 3D Waterfall 0°’.

## Slide 21
5. Peak fit of a laterally inhomogeneous sample
Exercise 5
Fit of XP spectrum of a laterally inhomogeneous sample with two different materials and loss structures.
Advanced peak fit using the Tougaard background for inhomogeneous samples.
Differences between different background models
1. Shirley background, 2. Tougaard background for homogeneous samples and 3. Tougaard background for inhomogeneous samples are shown.
Note: Number of Tougaard-background functions for inhomogeneous samples and number of peak-fit components have to be the same.
Exercise-05-Start two peaks, Voigt profile, Tougaard background, polynomial, statistic noise,
1. Peak: 20 000 counts, Lorentzian and Gaussian width: 3 eV, position: 100 eV, 2. Peak: 4,000 counts, Lorentzian and Gaussian width: 3 eV, position: 120 eV

## Slide 22
5. Peak fit of a laterally inhomogeneous sample
Task 5
Set the peak-fit settings to ‘Convolution’ and ‘Absolute’. Use two components and a fittable background.
Load test spectra ‘Exercise-05-Start.ufp’, (three spectra in three windows).
First Window peak fit with Shirley background
Reduce the spectral range (recommended values: 138 eV – 80 eV).
Define a fittable background with [Peak Fit → Fit Background XPS (HOM)], using the given parameters.
Generate a peak fit with two singlet peaks.
Set the start values GP-FWHM and LP-FWHM: 2 free, for both components.
Start the iteration with 22 iterations and 4 cycles.
Polynomial /Shirley |  | Tougaard | 
a = 900 | free | B = 0 | fix
b = 0.1 | free | C | 
c = 0.001 | free | C’ | 
d = 0 | fix | D | 
e = 0.001 | free | T0 | 

## Slide 23
5. Peak fit of a laterally inhomogeneous sample
Second Window peak fit with homogeneous Tougaard background
Set the background function to homogeneous Tougaard background with [Preferences – Preferences], ‘Tougaad Background Calculation’: Homogeneous Samples’.
Define a fittable background with [Peak Fit → Fit Background XPS (HOM)], using the given parameters (saved in Exercise-05-HOM.cro).
Generate a peak fit with two singlet peaks.
Set the start values GP-FWHM and LP-FWHM: 2 free, for both components.
Start the iteration with 22 iterations and 4 cycles.
polynomial/ Shirley |  | Tougaard | 
a = 900 | free | B = 2000 | free
b = 0.1 | free | C = 450 | free
c = 0.001 | free | C’ = 4.5 | free
d = 0 | fix | D = 500 | free
e = 0 | fix | T0 = 0 | fix

## Slide 24
5. Peak fit of a laterally inhomogeneous sample
Third Window peak fit with inhomogeneous Tougaard background
Set the background function to: inhomogeneous Tougaard background with ‘Tougaard Background Calculation’: Inhomogeneous Samples.
Define a fittable background with [Peak Fit → Fit Background XPS (INHOM)] and set the number of Tougaard background functions to 2.
Use the given parameters (saved in Exercise-05-P1.cro and Exercise-05-P2.cro).
Generate a peak fit with two components and singlet peaks.
Set the start values GP-FWHM and LP-FWHM: 2 free, for both components.
Start the iteration with 4 iterations and 4 cycles.
Polynomial/Shirley |  | Tougaard | Peak 1 | Tougaard | Peak 2
a = 900 | free | B = 200 | free | B = 330 | free
b = 0.1 | free | C = -340 | free | C = 550 | free
c = 0.001 | free | C’ = 2.8 | free | C’ = -0.9 | free
d = 0 | fix | D = 500 | free | D = 550 | free
e = 0 | fix | T0 = 0 | fix | T0 = 0 | fix

## Slide 25
5. Peak fit of a laterally inhomogeneous sample
3D-presentation of the background functions
Activate ‘Batch Processing → Plot 3D Waterfall 0°’, select all windows and leave dialogue with OK.
The dialogue ‘Curve Selection’ will open, select ‘Background’ and leave the dialogue with OK.
Design the 3D waterfall plot with the ‘Annotation/Design’ functions.
Open the design functions for the X-axis, Y- axis, curves and coordinate system with the right mouse key at the correct position.

## Slide 26
6. Calculation of a transmission function T(E)
Exercise 6
Calculation of a transmission function from a survey spectrum.
Verification of the accuracy of the saved T(E) functions.
Comparison with a quantification via Wagner factors.
Exercise-06-Start real data of Cu for the analysis of the T(E) function, Al-twin anode, LAXL lense, pass energy: 50 eV and 10 eV

## Slide 27
6. Calculation of a transmission function T(E)
Task 6
Set ‘the background function to homogeneous Tougaard background.
Load ‘Exercise-06-Start.ufp’ and maximize the survey spectrum window
Calculate the transmission function from the survey spectrum via ‘Calibration Intensity Scale → Load Reference Spectrum → Cu-twin’.
Calculate the fit parameters via iteration of a0, b1 and b2 using 200 iterations and 100 cycles.
Save the calculated T(E) in ‘Documents\Unifit_2024_User_Files\XPS-transfct’.
Quantification of Cu 2p and Cu 3p lines, 50 eV pass energy, with
a) Wagner factors and b) Sigma, Lambda and the corresponding T(E).
Quantification of Cu 2p3 and Cu 3p lines, 10 eV pass energy, with
a) Wagner factors and b) Sigma, Lambda and the corresponding T(E).
Compare the results of the quantifications.
Plot the estimated T(E) with ‘Calibration Intensity Scale → Plot Transmission function’ (Excitation energy: 1486.6 eV, Initial energy: 200 eV, Final energy: 1500 eV).

## Slide 28
7. Qualitative and quantitative analysis of XPS-standard measurements
Exercise 7
Spectra modification, e.g. charge correction, …
Identification of lines using different options.
Identification of chemical states.
Labeling of spectra.
Exercise-07-Start survey spectrum and C 1s, N 1s, O 1s and S 2p detail scans

## Slide 29
7. Qualitative and quantitative analysis of XPS-standard measurements
Task 7
Set preferences to ‘Sum’, ‘Relative’ and ‘Homogeneous Samples’.
Load ‘Exercise-07-Start.ufp’.
Activate C 1s and find the energy of the peak maximum with ‘Information → Minimum/Maximum’.
Reference the binding energies of all peaks to CH2 = 285 eV via ‘Batch Processing → Charge Correction All Windows’.
Identify and label all peaks of the survey spectrum with ‘Information – Identify Lines’ and the options ‘Important Lines’ and ’Element Lines’ and ‘Manual’.
Perform the satellite subtraction of C 1s, O 1s, N 1s, S 2p and label all spectra with ‘Annotation/Design – Spectrum Title 1’.

## Slide 30
7. Qualitative and quantitative analysis of XPS-standard measurements
Task 7
Perform the fit of the C 1s spectrum (three Peaks), O 1s spectrum (two Peaks) and N 1s spectrum (two Peaks).
(Keep L-G Mixing and FWHM the same for all lines of one peak. Background parameters a, b, c and e free.)
Fit the S 2p with two doublets and the given parameters.
Annotation of the peak-fit components with ‘Information – Identify Lines’, option: Chemical Shift - S 2p: -S-, S6+, C 1s: -CH2-, -C-O-R, N 1s: -N-CO-, -NH3+, O 1s: >C=O, -C-O-R
Open the quantification table via ‘Quantification → Concentration’.
 | doublet 1 | doublet 2
intensity ratio | 2:1 fix | 2:1 fix
L-G Mixing | 1 free, 1 fix | 1 fix, 1 fix
position second peak | 1.2 fix | 1.2 fix
FWHM | 1 free, 1 fix | 1 free, 1 fix

## Slide 31
8. Qualitative and quantitative analysis of AES-standard measurements
Exercise 8
Identification of lines.
Labeling of spectra.
Quantification of differentiated spectra, using empirical sensitivity factors.
Exercise-08-Start AES survey spectrum

## Slide 32
8. Qualitative and quantitative analysis of AES-standard measurements
Task 8
Load ‘Exercise-08-Start.ufp’.
Set ‘Preferences → Points to Average’ to 3.
Copy the spectrum in four additional windows and generate the derivative in those four windows (Setting: All Std. Windows).
Identify and label all peaks of the integrated spectrum (Ti (LMM), O (KLL), N (KVV), C (KLL)).
Reduce the energy range of the differentiated spectra window Ti 396-436 eV, O 497-537 eV, N 349-397 eV, C 230-285 eV.
Give every window the correct peak name by ‘Annotation/Design → Edit Parameters Standard Windows → Peak name’.
Open ‘Quantification → Concentration’ and select the differentiated windows.
Load the empirical SFs ‘Wagner-AES-PHI700-3kV-Differntial.sen’ and press ‘Calculate’.

## Slide 33
9. Layer thickness estimation
Exercise 9
Performance of a quantification.
Estimation of the layer thickness via energy resolved XPS (ERXPS) of an uncovered and a covered sample.
Exercise-09A-Start Cu 2p3/2 and Cu 3p detail scans of an covered and of an uncovered copper Exercise-09B-Start sample, twin anode, lense mode: SAE150, Pass energy: 50 eV

## Slide 34
9. Layer thickness estimation
Task 9A Layer Thickness Estimation by energy-resolved XPS (ERXPS)
Activate the preference option ‘Open Projects with Showing of the Quantification and Film Thickness Table’ (ticked).
Set number of the first standard window to 11 via ‘Preferences → Preferences → Programme Parameters → Number of the First Standard Window’
Load the file ‘Exercise-09A-Start.ufp’ (uncovered sample).
Table of concentration pops up, close it with ‘Cancel’.
Choose ‘Quantification → Layer Thickness Estimation 1 (ERXPS)’.
Choose the Cu 2p3/2 and Cu 3p lines (thickness should be 0 (dashed lines))
Close all windows and load the file ‘Exercise-09B-Start.ufp’ (covered sample).
Open the quantification table via ‘Quantification – Concentration’ and chose
Cu 2p3/2 and Cu 3p. Close the quantification table with ‘Cancel’.
Choose ‘Quantification → Layer Thickness Estimation 1 (ERXPS)’.
Choose peaks Cu 2p3/2 and Cu 3p (calculated thickness about 10 Å).

## Slide 35
9. Layer thickness estimation
Task 9B Layer Thickness Estimation by angular-resolved XPS (ARXPS)
Load the file ‘Exercise-09A-Start.ufp’ (uncovered sample).
Table of concentration pops up, close it with ‘Cancel’.
Choose ‘Quantification → Layer Thickness Estimation 2 (ARXPS) → Calculation/Input of Calibration Values…I’(x°)(d=0)’.
Choose the Cu 2p3/2 and save the intensity for d=0 by using the ‘Save’ button the folder ‘Unifit_2024_User_Files/Film’ and the name is ‘Cu2p3_uncovered.dat’.
Close all windows and load the file ‘Exercise-09B-Start.ufp’ (covered sample).
Open the quantification table via ‘Quantification – Concentration’ and chose
Cu 2p3/2. Close the quantification table with ‘Cancel’.
Choose ‘Quantification → Layer Thickness Estimation 2 (ARXPS) → Calculation of Thickness d with I’(x°; d=0)/I’(x°;d>0)’.
Choose Cu 2p3/2 and load the saved intensity for d=0 by ‘Load I’(x°) d=0’’ (calculated thickness about 11 Å).
Compare the calculated values.
(In case you are ready, you can run the same procedure with Cu 3p.)

## Slide 36
Exercise 10
Batch processing of data series (spectra depending on one parameter, e.g. angle, sputter time, time of deposition, temperature, …).
3D and parameter plot of the series.
Design of the parameter plot window.
Note: Not in all cases a complete fit of the data is necessary, in order to get good quantification results.
10. Batch processing
Exercise-10-Start naturally oxidized Si (→ SiO2)

## Slide 37
Task 10
Set preferences to ‘Product’, ‘Absolute’ and ’Homogeneous Samples’. First Standard Window: 11.
Open ‘Exercise-10-Start.ufp’. All Si 2p peaks are plotted. Activate window 11.
Si 2p peaks
Run the satellite subtraction, an appropriate reduction (111 eV – 91 eV) and the peak fit with two doublets (fix the parameters of the second peak of the doublet) and use the ‘Fit Background XPS (HOM)’ with the parameter setting: (a, b, e free, all others zero and fixed) in window 11.
Run the batch processing of all Si peaks via ‘Batch processing → Batch Processing’.
Calculate the concentration with the correct transmission function (Al-twin, 10 eV Pass, SAE 150) and sensitivity factors (AlScofld.sen).
Create the parameter plot. (x axis: Batch parameter, y axis: Norm. Areas (Components))
Close the concentration table and all standard windows (parameter plot still open).
10. Batch processing

## Slide 38
Task 10
O 1s peaks
Open all O 1s peaks via ‘File → Select Blocks’.
Appropriately subtract the background and run the batch processing by activation of ‘Setting’ All Std. Windows’.
Calculate the concentration and add the results to parameter plot of Si 2p.
Close the conc. table and all standard windows (parameter plot still open).
C 1s peaks
Open all C 1s peaks via block selection, subtract the background and run the batch processing by activation of ‘Setting’ All Std. Windows’.
Calculate the concentration and add the results to the parameter plot.
Design the parameter plot via ‘Preferences → Display → Parameter Plot Display’ or use the right mouse key.
Present Si 2p spectra with different 3D plots (e.g. ‘Batch Processing → Plot 3D Colour Profile’).
10. Batch processing

## Slide 39
11. Processing of a line scan
Exercise 11
Estimation of the recording size of a specific spectrometer setting.
Batch processing of a line scan over a sample edge.
Exercise-11-Start Si 2p signal line scan of a sputtered depth profile of naturally oxidized Si (→ SiO2), batch parameter includes the manipulator x-position, scaling: micrometer

## Slide 40
11. Processing of a line scan
Task 11
Set preferences to ‘Product’ and ‘Absolute’, number of the first window: 11, Maximal Number of Presentable Steps inside ‘Parameter Plot’: 200.
Open ‘Exercise-11-Start.ufp’ and display every 8th standard window, using ‘Windows → Hide Standard Spectra’.
Activate window 32, define a fittable background (a free, b free, c = 0 fix, d = 0 fix, e free, B = 0 fix) and carry out the peak fit, using two singlets: Si at 99.6 eV and SiO2 at 103.6 eV.
Run the batch processing of all Si 2p peaks. Attention Before batch processing is started, fix the peak positions of the two components of window 32 and run the peak fit again!
Create a parameter plot of the fitted areas of all windows with ‘Batch Processing → Plot Fit Parameters/Min/Max/Area’ (x axis: Batch parameter, y axis: abs. Area).
Remove the plot of the SiO2 component from the parameter plot using ‘Annotation/Design → Legend/Select Curves’ (parameter plot activated).
Plot only the first curve and change the name of the first curve to ‘Si 2p’.
Change the position of the legend to ‘right’ and ‘above’.
Design the parameter plot for a good presentation.
Estimate the recording size.

## Slide 41
Exercise 12
Batch processing of a large number of spectra using hidden standard windows.
Dynamic illustration of a sputter process (size and position of the sputter crater during the sputter process).
12. Processing of a sputter depth profile multipoint measurement
Exercise-12-Start GaAs sample, As 2p3/2 signal multipoint measurements of a sputter depth profile, (11x9 measurement points, 10 sputter steps → 990 As 2p3/2 spectra)

## Slide 42
12. Processing of a sputter depth profile multipoint measurement
Task 12
Set number of the first standard window to 41.
Set the preferences to ‘Product’ and ‘Absolute’.
Display every 128th standard window.
Open ‘Exercise-12-Start.ufp’.
Change the x and y positions with ‘Information/Editing → Edit Parameters Standard Windows’ and the operation ‘All Windows’ and ‘Fill with’ to x-Position from -3 to 3 mm (division by 1000, subtraction with 27.5) y-Position from -3.75 to 3.75 mm (division by 1000, addition with 2.10).
Activate window 1024, load the saved fit parameter data ‘Exercise-12-Start.par’ (two singlet peaks, fittable background) with ‘Peak Fit → Load Start Parameters’ and confirm the parameters with OK.
first component: GaAs species (bulk), second component: As oxide species (surface)
Perform the peak fit with 10 iterations and 2 cycles.
Start the batch processing.

## Slide 43
12. Processing of a sputter depth profile multipoint measurement
Task 12
Generate ten ‘XY 3D Colour Profile’ of the peak area of the first component (bulk) with ‘Batch Processing → XY 3D Colour Profile’, select the first 99 standard windows (41 – 139), select ‘Area Component 1’ and ‘Number of 3D Plots’ is 10.
Start the generation of the ten 3D plots with OK (in the windows 1 – 10).
Repeat the generation of the ‘XY 3D Colour Plots’ with the surface species (Area Component 2), plotted in the windows 11 – 20.
Generate additionally the ‘XY 3D Plot 45’ of the bulk and surface component in the windows 21 -30 and 31 – 40 and edit them to your satisfaction.
Activate the video sequences via ‘Windows → Show Windows Video Sequence’. They should include the windows 1 – 10, 11 – 20, 21 – 30 and 31 – 40.
Set the frame-dwell time to 2 sec.
Activate ‘Same Position of Windows as Active Window’, ‘Same Window Size as Active Window’ and ‘Same Scaling of Intensity Axis of 3D Window’.
Start the video sequence with OK.

## Slide 44
13. Batch Processing of Multi-Region Measurements
Exercise 13
Loading of multi-region measurements.
Generation of a processing-master, with appropriately fixed background parameters, fit parameters and fit-parameter limits.
Saving of processing and design steps of all regions of one measurement (master) into a *.ppd file.
Reloading and applying of processing-master steps for all loaded spectra, common quantification of samples.
Exercise-13-Start Survey, O 1s, C 1s, Ga 3d, As 3d, Ga 2p3/2, As 2p3/2, VB spectra of three measurements, processing steps: charge correction, spectrum title

## Slide 45
13. Batch Processing of Multi-Region Measurements
Task 13
Set preferences to ’Sum’ and ‘Relative’, the x axis to ‘XPS: Binding Energy/XAS: Photon Energy’ and the first standard window to 11, show all standard windows.
Load ‘Exercise-13-Master.ufp’, have a look at the fit parameters as well the fit-parameter limits, the annotations and the design.
Save the processing steps of all windows with ‘File → Save Project Processing Steps/Design Standard Windows’, use the file name: ‘My Unifit Project Processing Steps\Master.ppd’, close all windows.
Load the multi-region measurements of three GaAs samples using the batch file loading, open the first data set with ‘File → Open Spectra → XPS → VAMAS’ from ‘Exercise-13-Start\GaAs-01.vms’, all files with the same name and an appropriate batch name (-01, -02, -03) are loaded (option: Internal Spectra Names: Peak name plus _Batch parameter), show all standard windows.
Entitle all spectra with ‘Sample’ + batch parameter by - activate window 11, open ‘Annotation/Design → Spectrum Title 1’, - type the name ‘Sample’ and check ‘All Windows of the Same Type’ and check ‘Title + Batch Parameter’.

## Slide 46
13. Batch Processing of Multi-Region Measurements
Task 13
Carry out the charge correction for the regions of sample 1, by - activate the C 1s spectrum, determine the energy of the peak maximum and calculate the difference to 285.0 eV (1.5 eV), - define the charge correction of the regions of sample 1 with ‘Batch Processing → Charge Correction…’, check all regions of sample 1 (window 11 - 18) and press OK.
Repeat the charge correction for the regions of sample 2 (0.35 eV) and sample 3 (0.75 eV), check the charge correction with ‘Information → Charge Correction/Wavenumber Offset’.
Load the previously saved processing-master steps with ‘File → Open Project Processing Steps/Design Standard Windows’, file name: ‘My Unifit Project Processing Steps\Master.ppd’
The windows-selection dialogue appears, check all windows (11 - 34) and press OK.

## Slide 47
13. Batch Processing of Multi-Region Measurements
Task 13
An opening message box gives information about the saved processing-step names and the title annotation. Press OK to start the batch processing.
Check all windows of correctness of the peak fit and annotation as well as the annotation positions: - start additional iterations of window 27 and window 30 for a better results. - create a new fit of windows 32 with two fit components, since the O 1s signal of sample 3 is different with respect to sample 1 and 2. Uncheck the fixed FWHM of O 1s of the first component before iteration.
Carry out the quantification of the results using the Ga 3d, As 3d, C 1s and O 1s lines
- via ‘Quantification → Concentration’, - select the windows of interest (12, 13, 15, 16; 20, 21, 23, 24; 28, 29, 31, 32), - press ‘Par. Lines 100%’ to set the 100% normalization separately for each sample.

## Slide 48
14. Analysis of a Micro-RAMAN measurement of solar silicon
Exercise 14
Micro-RAMAN measurement (51x51 spectra) of solar-silicon wafer.
Analysis of the defect intensity (FWHM), internal stress (peak intensity), doping (asymmetry) and grain orientation (intensity).
Comparison of the results after the peak fit and the results using the original measurement data.
Task 14
Setting of all general programme parameters to minimum but: - Maximal numbers of presentable curves inside ‘Plot 3D Waterfall’: 300 - Maximum number of presentable curves inside ‘Parameter plot’: 30 - Number of the first standard window: 41
Load the Unifit project ‘Exercise-14-Result.ufp’.
Evaluate and discuss the results presented in the 3D presentations.
Generate a XY 3D Color Profile of the positions of the maxima of the original measurement data.
Compare the peak fit results with the results of the original measurement data.

## Slide 49
15. Export presentation of results
Exercise 15
Export of fitted spectra via ‘Copy Image’ and ‘Paste’.
Export of fitted spectra as *.emf (400 dpi).
Export of concentration table via ‘Copy – Paste’.
Export of fit parameter table via ‘Copy – Paste’.
Paste into a MS Word document.
Task 15
Load ‘Presetting_Image_Export.set’ via ‘Preferences → Load Preferences…’.
Open ‘Exercise-04B-Result-All.ufp’.
Change resolution ‘Preferences → Resolution Image Export/Copy’ to 400 dpi.
Save image in clipboard via ‘File → Copy Image’.
Start a new MS Word document, import image via ‘Modify → Copy’.